## Slide 1: Pulsed Fiber Laser Autocorrelator
Matthew Wilson and Gabe Trippel

## Slide 2: What Is Optical Autocorrelation?
A way to measure a femtosecond laser pulse
Uses a second-order response (nonlinear) amplified by constructive interference
Defined by convolution of electric field of two different paths
A(t) ≡ ʃ E(t)E*(t - Δt) dt
+
=
(constructive part only)

## Slide 3: How it works…
Uses Michelson interferometer
  Laser pulses constructively interfere when Michelson arm lengths are equal
Requires second order response detection
  P(t) = ɛ0(χ(1)E(t) + χ(2)E2(t) + …) (polarization density)
  Achieved by…
    SHG with nonlinear crystal
  …or…
    Two-photon absorption
Amplitude of nonlinear signal is proportional to intensity:

## Slide 4: Pulsed Fiber Laser System
From Wesley Hughes and Jared Green’s Presentation

## Slide 5: Signal Detection System
Observed Nonlinear Response of IR Fiber Laser:
  For Both Standard Silicon Diode (FDS100, ThorLabs) and Avalanche Photo Diode (AD500, Pacific Silicon Sensors)
  Requires small beam diameter on detector to increase probability of Two Photon Absorption
    d = λf/∏D D = Diameter of Beam d = diameter of focused spot size
    f = focal length
Also obtained the Gain Curves for :
  532nm Diode Laser, HeNe Laser, and IR Pulsed Fiber Laser

## Slide 6: Gain of Avalanche Photo Diode
Gain = (APD Output) / (ADP Output at 1V)
Chart: Gain Curve (Linear Response)
| 0.999 | 10.050000000000002 | 20.03 | 30 | 40.01 | 50.04 | 60.02 | 70.099999999999994 | 80 | 90 | 100 | 110 | 120 | 130 | 140 | 150
HeNe Laser (632 nm) | 1 | 1.04 | 1.1 | 1.63 | 3.37 | 5.1 | 6.12 | 7.39 | 8.98 | 10.92 | 14.29 | 18.92 | 26.94 | 40.82 | 63.27 | 94.29
Diode Laser (532 nm) | 1 | 1.09 | 1.1 | 1.2 | 2.08 | 2.8 | 3.3 | 4 | 4.86 | 5.9 | 7.38 | 10 | 15.1 | 26 | 46 | 79

## Slide 7: APD Output vs Bias Voltage for Pulsed Fiber Laser
Chart: APD Output vs Bias Voltage
| 90.1 | 100 | 110 | 120 | 130.4 | 133 | 134 | 135 | 136 | 137 | 138 | 139 | 140
IR Laser | 2.53 | 2.41 | 2.7 | 2.69 | 2.44 | 2.4 | 2.45 | 2 | 4 | 26 | 55 | 44.8 | 60
No signal below 90V Reversed Bias
No External Amplification on APD (yet)

## Slide 8: Interferometer Diagram and Picture
Concave Lens
Convex Lens’s
Laser
Mirror
Mirror
Mirror
Mirror
Speaker w/ Corner Cube
Mirrors and Lenses are maximized for IR
Beam Splitter
Beam Chopper included between Beam Splitter and Speaker
Top and Bottom Arm’s are chopped at different frequencies
Chopper

## Slide 9: Data Acquisition
Linear Response measured with InGaAs Detector and Continuous Wave 1550nm IR Diode Laser
Speaker Off Speaker On (1 Hz)

## Slide 10: Nonlinear Response!
Nonlinear Response from 1550 nm IR Erbium Doped Pulsed Fiber Laser
Beam Passed through Interferometer then Focused onto Avalanche Photo Diode (APD)
  Gain: G ≈ 25 internal to APD at 135V, plus external gain from LT1028 op-amp
  Beam Blocked Beam Unblocked
Beam Blocked Beam Un-Blocked
Response Visible on APD Output (Lock-In Amp not able to Lock onto signal)

## Slide 11: Future Investment
What can help to Complete Project:
InGaAs and Silicon (layered) Detector
  Allows Observation of Linear & Nonlinear Response Simultaneously (for easier signal capture/maximization)
Automatic single-axis translational stage for one Michelson arm
  Because human arms get tired
XYZ translation mount for Detector
  For Signal Optimization
More Powerful Laser

## Slide 12: New techniques for laser pulse measurement:
FROG – Frequency Resolved Optical Gating
  Similar to autocorrelation, but gets more info
    Asymmetries resolved (chirp, satellite pulses)
    FROG trace:
SPIDER – Spectral Phase Interferometry for Direct Electric field Reconstruction
  Uses spectral shearing
  Both techniques require a spectrometer and special software.
  (Other “swamp” techniques: BOA Compression, SEA TADPOLE, MIIPS, and more…)

## Slide 13: This is The End……